## Slide 1
Few-Shot Personalized PPG-Based
Blood Pressure Estimation
Progress Update · Cheng Jiangyu · 17 Aug 2026
Notes: [Sources]
- Internal project progress report; no external claims on this slide.
[/Sources]

## Slide 2
Leakage-Safe Few-Shot Evaluation Protocol
2
Development data
PulseDB development cohort
Subject-disjoint meta-validation
Chronological event ordering
Calibration
K = 1 / 2 / 3 / 5 reference events
First K events are support
Predict event 6 and later queries
Primary evaluation
697 participants
103,564 common queries per K
Participant-macro MAE
Population PPG model
→
K-event calibration
→
Event 6+ prediction
Reference BP values in PulseDB are ABP-derived labels, not literal cuff readings.
Meta-validation only · Locked meta-test untouched
Notes: [Sources]
- C:\Research\Calibration PPG BP\docs\STATUS.md
- C:\Research\Calibration PPG BP\docs\METHODS.md
[/Sources]
The locked meta-test is not accessed during model selection.

## Slide 3
Models and Personalization Mechanisms
3
Setting
Core mechanism
Personalization behavior
B0
Population BP mean
No PPG; no calibration
B1
Last reference BP
Copies the latest support label
B2
Mean support BP
Averages K support labels
B3
Population PPG network
PPG-only; no personal adaptation
B4
Residual offset
Adds mean support residual to B3
B5
Siamese metric model
Support-query similarity; K=1 diagnostic
B6
Head-only adaptation
Updates only the prediction head
B7
Full fine-tuning
Updates the complete network
B8
LoRA adaptation
Low-rank parameter update
M0
Residual-anchor set encoder
Mean residual anchor + support context
M1
M0 + FiLM
Feature-wise modulation
M2
M0 + reliability weighting
Query-conditioned support weighting
M0 is the current parsimonious candidate; M1/M2 add complexity for ablation.
Notes: [Sources]
- C:\Research\Calibration PPG BP\docs\METHODS.md
[/Sources]

## Slide 4
Training Round 1 — Single-Seed Exploration
4
Setting
K=1
K=2
K=3
K=5
Interpretation
B0 Population mean
13.620
13.620
13.620
13.620
No personalization
B1 Last reference
15.023
13.825
12.993
12.733
Label-only
B2 Support mean
15.023
13.570
12.755
11.948
Label-only
B3 Population PPG
11.638
11.638
11.638
11.638
PPG-only
B4 Residual offset
12.305
11.190
10.593
9.931
Simple calibration
B6 Head-only
10.975
10.427
10.040
9.554
Gradient adaptation
B7 Full fine-tune
15.032
13.381
12.498
11.862
Unstable few-shot fit
B8 LoRA
10.372
10.043
9.824
9.509
Parameter-efficient
M0
9.998
9.435
9.119
8.688
Best overall
M1
10.307
9.630
9.280
8.818
M0 + FiLM
M2
10.622
9.807
9.381
8.873
M0 + weighting
Participant-macro mean MAE (mmHg) · seed 20260813
Meta-validation only · Locked meta-test untouched
Notes: [Sources]
- C:\Research\Calibration PPG BP\docs\RESULTS_PHASE5.md
[/Sources]

## Slide 5
Training Round 2 — Five-Seed Early-Stopping Validation
5
Setting
K=1
K=2
K=3
K=5
Change vs Round 1
B0 Population mean
13.620
13.620
13.620
13.620
—
B1 Last reference
15.023
13.825
12.993
12.733
—
B2 Support mean
15.023
13.570
12.755
11.948
—
B3 Population PPG
11.710
11.710
11.710
11.710
+0.072
B4 Residual offset
12.327
11.200
10.624
9.951
+0.020
B6 Head-only
10.963
10.472
10.105
9.588
+0.034
B7 Full fine-tune
15.042
13.382
12.520
11.842
−0.020
B8 LoRA
10.499
10.140
9.944
9.608
+0.099
M0
10.081
9.526
9.235
8.785
+0.097
M1
10.190
9.563
9.251
8.807
−0.011
M2
10.233
9.641
9.332
8.834
−0.039
Participant-macro mean MAE (mmHg) · mean across 5 seeds
Meta-validation only · Locked meta-test untouched
Notes: [Sources]
- C:\Research\Calibration PPG BP\docs\RESULTS_PHASE5_REPEAT5.md
[/Sources]

## Slide 6
M0 Performance by Calibration Budget
6
Calibration
Overall MAE
SBP MAE
DBP MAE
Gain vs B3
K=1
10.081
12.920
7.242
1.629
K=2
9.526
12.240
6.812
2.184
K=3
9.235
11.894
6.576
2.475
K=5
8.785
11.318
6.253
2.925
9.407 ± 0.104
Four-K mean MAE across 5 seeds
0.046 mmHg
M0 advantage over M1 (four-K mean)
0.103 mmHg
M0 advantage over M2 (four-K mean)
M0 remains the simplest leading model; small M0–M1/M2 gaps are not claimed as decisive.
Notes: [Sources]
- C:\Research\Calibration PPG BP\docs\RESULTS_PHASE5_REPEAT5.md
[/Sources]

## Slide 7
Historical Calibration-Free Context
7
Experiment
Dataset / protocol
SBP MAE
DBP MAE
Interpretation
Historical XResNet-50
VitalDB; calibration-free; window-level
12.492
8.059
Prior baseline
Current M0 (K=5)
PulseDB; few-shot; participant-macro
11.318
6.253
Personalized model
Not an apples-to-apples comparison
The dataset subset, split, metric aggregation, and calibration protocol differ. The historical result is shown only to motivate the transition from calibration-free to personalized prediction.
Meta-validation only · Locked meta-test untouched
Notes: [Sources]
- Prior VitalDB calibration-free XResNet-50 experiment summary
- C:\Research\Calibration PPG BP\docs\RESULTS_PHASE5_REPEAT5.md
[/Sources]
Do not interpret the numeric difference as a controlled head-to-head gain.

## Slide 8
Worst 20% Subjects Show Calibration Drift
8
Indicator (K=5)
Remaining 80%
Worst 20%
Observed pattern
Participant mean MAE
6.935
15.662
2.26× higher
Support→query SBP change
13.704
22.660
+65%
Support→query DBP change
7.288
12.791
+76%
Within-subject SBP / DBP SD
12.056 / 6.679
16.011 / 9.426
+33% / +41%
Query outside support BP range
55.9% / 54.4%
70.4% / 69.4%
More extrapolation
Mean event index
72.5
108.7
+50% later
Filtered-PPG SD / mean age
0.2753 / 61.8
0.2748 / 63.1
Nearly unchanged
26.6% MIMIC
Worst-tail participant rate
14.7% VitalDB
Worst-tail participant rate
140 / 697 participants · threshold ≥ 11.751 mmHg · diagnostic association, not proven causality
Notes: [Sources]
- C:\Research\Calibration PPG BP\docs\PHASE6_HIGH_ERROR_ANALYSIS.md
[/Sources]
The worst tail is defined using observed meta-validation error and is therefore oracle-only. PPG SD is an amplitude proxy, not a validated signal-quality index.

## Slide 9
Phase 6 Tests One Change at a Time
9
Candidate
Only change from M0
Purpose
Job
Robust loss
Huber loss
Reduce large-residual influence
819
BP-change sampling
Meta-train sampler
Learn calibration-drift episodes
822
Robust anchor
Median residual anchor
Resist atypical cuff anchors
823
PPG quality gate
PPG-only gate
Attenuate unreliable personalization
824
Age + sex
Clean demographic vector
Add subject context
825
Fixed-first job 818 complete: four-K mean 9.137 mmHg (−0.173 vs same-seed M0); K1–3 improved, K5 +0.094
Age uses participant median; invalid age is retained with a mask; normalization uses meta-train only.
Notes: [Sources]
- C:\Research\Calibration PPG BP\docs\PHASE6_EXPLORATION.md
- C:\Research\Calibration PPG BP\docs\PHASE6_HIGH_ERROR_ANALYSIS.md
[/Sources]
Jobs 819 and 822–825 are isolated single-factor candidates. Fixed-first job 818 is a single-seed protocol ablation, not a final result. Locked meta-test remains untouched.